# CRSSIS Microtops Sun Photometer II Guide.

Tony Gill and Stuart Phinn

Based on the CSIRO Microtops II Quick Guide and the microtops II User’s Guide

Overview:

The sun photometer is used for measuring the aerosol optical thickness (or aerosol optical depth) of the atmosphere. The sun photometer is also used to measure the irradiance (amount of light reaching the ground) in 5 optical bands (440 nm, 500 nm, 870 nm, 936 nm, 1020 nm). Optical depth is described as (From http://en.wikipedia.org/wiki/Optical_depth)  a measure of transparency, and is defined as the negative logarithm of the fraction of radiation (or light) that is scattered or absorbed on a path. One way of visualizing optical depth is to think of a fog. The fog between you and an object that is immediately in front of you has an optical depth of zero. As the object moves away, the optical depth increases until it reaches a large value and the object is no longer visible.

The optical depth expresses the quantity of light removed from a beam by scattering or absorption during its path through a medium. If I0 is the intensity of radiation at the source and I is the observed intensity after a given path, then optical depth τ is defined by the following equation:

In atmospheric sciences, one often refers to the optical depth of the atmosphere as corresponding to the vertical path from Earth's surface to outer space; at other times the optical path is from the observer's altitude to outer space. Since τ refers to a vertical path, the optical depth for a slant path is τ' = mτ, where m is called the airmass factor, and for a plane-parallel atmosphere it is determined as m = 1 / cosθ, where θ is the zenith angle corresponding to the given path. Therefore

Contents of the system

- Microtops II sun photometer

- Black carry bag/case (the sun photometer must remain in this bag at all times).

- Grey data cable

- User’s guide

- CD with software

- (optional) when going in the field take an extra set of 4 AA alkaline batteries. A new set of batteries will last at least 50 hours of continuous operation. AA NiCd will not provide sufficent voltage!

Starting and configuring the system

- Press the ON/OFF button to turn on the instrument. Keep the channel window on top of the instrument covered. Always be careful with this window. Keep it closed when not measuring and do not touch.

- Instrument then displays “Hardware Test”, Firmware version and starts the initialisation.

- Delete any measurements that are in memory from another user (they should have downloaded this data already!)

  - Press menu/enter

  - Use ↓ to scroll down to Data Logging and press menu/enter

  - Press menu/enter on Clear Memory

  - Press right arrow key to clear all.

Before taking measurements there are values that must be set:

- Location: Latitude and Longitude measurement from GPS (must be WGS84). To set:

  - Press menu/enter

  - Scroll down to Location and press menu/enter

  - Scroll down to Coordinates and press menu/enter

  - Use the <- and -> arrows to move cursor to value to be changed. Use ↑or ↓ to change the values.

  - Press menu/enter when done.

  - Press scan/escape twice to return to main menu.

- Altitude: altitude from the GPS (although this is not used in computations for sun photometer, it is for documentation only). To set:

  - In the Location menu, scroll down to Altitude and press menu/enter.

  - Set the altitude using the arrow keys and press menu/enter to set.

  - Escape back to main menu.

- Pressure: The sun photometer has an in-built pressure sensor. The pressure value must be set to zero (default) for the pressure sensor to be used, otherwise the user specified value is used. To set (or check that it is zero):

  - In the Location menu, scroll down to Pressure and press menu/enter.

  - Set the pressure using the arrow keys and press menu/enter to set.

  - Escape back to main menu.

- Time: Time must be universal time (not local time) and match the GPS. The time on the system will be close to the GPS time, but should be checked and reset if necessary. To set the time:

  - Press menu/enter

  - Press menu/enter on the Clock option

  - Press menu/enter on the Adjust Clock option

  - Use the arrow keys to adjust the time.

- ID: displays on the main screen. This can be changed for each measurement by pressing the ↑or ↓ key when in ready  (RDY) mode.

- Scan length: number of scans per measurement. The number of scans depends on how stable the instrument can be held. Recommended value is 30, but if you are on a boat or other moving platform a value of 1 may be the most practical. To set:

  - Main menu->measurement->processing->scan length.

There are also some constants that are used in the calculations and will be provided when you download your data. DO NOT change these:

- C1-C5: Calibration constants for irradiance.

- LNV01-LNV05:  calibration constants for AOT

- K,B,C – calibration constants for water vapour

- POFFS, PSCALE – calibration constants for pressure

If these are changed accidently, they can be changed back to factory settings.

Taking a measurement

- Make a note of the weather conditions. Note that measurements under cirrus conditions are unusable.

- Have the sun photometer in the conditions in which measurements will be taken for at least 15 minutes before taking measurements to improve the temperature measurements (e.g. wait 15 minutes if the sun photometer has been in an air conditioned room).

- Open the top window and point the top of the instrument towards the sun. Try to catch the collimated sun beam in the bull’s eye.

- Keep the collimated beam centred and press the scan/escape button to start a measurement. Keep the beam as centred as possible. Measurements performed within the second or third circle are unusable.

- The instrument displays the scan number and beeps twice once a scan is completed. All scans are stored in memory. After switching On/Off new measurements are appended.

- When finished, switch off the instrument.

Downloading the data

- Install the Mtporg software from the CD. The registration information is:

  - Registered to: SN12597

  - License Key: ECB965DE3900

- Connect grey cable to the PC and instrument

- Switch instrument on

- Start Mtporg software

- Press the download button

- Enter a data description tag which will be added to the downloaded data base in the DATA_DESCR column.

- Press OK.

- The data is stored in a dbase file called DATA.dbf located in C:\Program Files\Solar Light Co\Microtops II Organizer.

- Close the download window

- Switch off the instrument

Viewing the data in the DBF file

A DBF file with more information than you need is provided as output – there seems to be a standard output for both the sun photometer and ozonometer. We are interested in the sun photometer output only. The DBF file can be opened and viewed in Microsoft excel. The following table provides details of the columns provided in the DBF file. Normally it’s only the AOT measurements that we are interested in.

Note, at the bottom of the table there are values with ? next to them. If anyone works out what these refer to please inform Tony Gill – t.gill1@uq.edu.au.

| Code | Description | Units |
| --- | --- | --- |
| SN | 5 digit serial number for the instrument with leading zeros |  |
| DATE | Universal date of the measurement |  |
| TIME | Universal time of the measurement |  |
| DATA_DESCR | Text description of data set provided by user |  |
| LOCATION | Location in Latitude and Longitude (WGS84) to be set by user |  |
| WATER | Water vapour column | cm |
| AOT440 | AOT at 440 nm | unitless |
| AOT500 | AOT at 500 nm | unitless |
| AOT870 | AOT at 870 nm | unitless |
| AOT936 | AOT at 936 nm | unitless |
| AOT1020 | AOT at 1020 nm | Unitless |
| SIG440 | Signal measured at the 440 nm sensor | mV |
| SIG500 | Signal measured at the 500 nm sensor | mV |
| SIG870 | Signal measured at the 870 nm sensor | mV |
| SIG936 | Signal measured at the 936 nm sensor | mV |
| SIG1020 | Signal measured at the 1020 nm sensor | mV |
| STD440 | Standard deviation of the signal measured at the 440 nm sensor | mV |
| STD500 | Standard deviation of the signal measured at the 500 nm sensor | mV |
| STD870 | Standard deviation of the signal measured at the 870 nm sensor | mV |
| STD936 | Standard deviation of the signal measured at the 936 nm sensor | mV |
| STD1020 | Standard deviation of the signal measured at the 1020 nm sensor | mV |
| R440_500 | Ratio of signals between 440 nm and 500 nm bands | unitless |
| R500_870 | Ratio of signals between 500 nm and 870 nm bands | unitless |
| R870_936 | Ratio of signals between 870 nm and 936 nm bands | unitless |
| R936_1020 | Ratio of signals between 936 nm and 1020 nm bands | unitless |
| C1 | Calibration constant for Irradiance at 440 nm | W.m-2 |
| C2 | Calibration constant for Irradiance at 500 nm | W.m-2 |
| C3 | Calibration constant for Irradiance at 870 nm | W.m-2 |
| C4 | Calibration constant for Irradiance at 936 nm | W.m-2 |
| C5 | Calibration constant for Irradiance at 1020 nm | W.m-2 |
| LNV01 | Natural logarithm of AOT calibration constant for 440 nm band | Unitless |
| LNV02 | Natural logarithm of AOT calibration constant for 500 nm band | Unitless |
| LNV03 | Natural logarithm of AOT calibration constant for 870 nm band | Unitless |
| LNV04 | Natural logarithm of AOT calibration constant for 936 nm band | Unitless |
| LNV05 | Natural logarithm of AOT calibration constant for 1020 nm band | Unitless |
| K | Water vapour calibration coefficient dependent on spectral transmission of the channel 936 nm filter | ? |
| B | Water vapour calibration coefficient dependent on spectral transmission of the channel 936 nm filter | ? |
| C | Scaling factor between aerosol optical depth at 1020 or 870 and 936 nm | ? |
| POFFS | Pressure calibration constant | mV |
| PSCALE | Pressure calibration constant | mB/mV |
| LATITUDE | Latitude of the site | Degrees |
| LONGITUDE | Longitude of the site | Degrees |
| ALTITUDE | Altitude above sea level | m |
| PRESSURE | Pressure (either preset or measured from sensor) | millibar (mB) |
| SZA | Solar zenith angle | Degrees |
| TEMP | Temperature of the optical block inside the instrument | Degrees celsius |
| AM | Optical airmass – the path length relative to that at zenith. Computed as 1/cos(SZA) | unitless |
| SDCORR | Earth sun distance correction (multiplicative) | unitless |
| ID | User defined identification code |  |
| SIG300 | Ozonometer only |  |
| SIG305 | Ozonometer only |  |
| SIG312 | Ozonometer only |  |
| SIG940 | Ozonometer only |  |
| R300_305 | Ozonometer only |  |
| R305_312 | Ozonometer only |  |
| STD300_305 | Ozonometer only |  |
| STD305_312 | Ozonometer only |  |
| OZ300_305 | Ozonometer only |  |
| OZ305_312 | Ozonometer only |  |
| OZONE | Ozonometer only |  |
| A1 | ? |  |
| A2 | ? |  |
| B1 | ? |  |
| B2 | ? |  |
| L1 | ? |  |
| L2 | ? |  |
| OC | ? |  |